CHILD & FAMILY DEVELOPMENT PROGRAMS

# Sensory Screening Procedure

Purpose:

To ensure all Head Start children receive sensory screening (vision and hearing) in collaboration with parents, within 45 days of entry.

Procedure for Hearing Screenings

- Hearing screenings will be done by appropriately trained staff on the Pilot Audiometer.

- Center Managers will have primary responsibility for assuring that all children are screened within 45 days. If child's hearing has not been screened within 30 days, a plan will be developed to involve the family to assure completion of screening before the 45 day deadline.

- All referrals will be completed by Winter Break for those children enrolling in September.

- All returning children will receive a hearing screening.

- All hearing screening results and follow-up plans and activities will be documented in Child Plus using CP2b and in the child's file to communicate results and follow-up plans to families.

- If a referral is made it will be made to the child's primary care provider and follow up/treatment will be coordinated by the provider.  Staff will assist parents in obtaining appropriate follow up and treatment as needed.

Pilot Audiometer: The Pilot Audiometer hearing machines are generally shared by more than one center.  A rotating schedule between sites is recommended for using the machine and for calibration, which must be completed annually prior to use.  To calibrate send to:

| Starkey Labs: ATTN: Calibration 6700 Washington Ave. S. Eden Prairie, MN 55344 (800) 733-2636 ext. 2937 | Alternate Option: MSR Northwest 800-950-EARS/3277 |
| --- | --- |
| Need PO – Reference model and serial number of machine and account #: E0079 | Need PO – Reference model and serial number of machine and account #: 01-0002795 |
| Package and mail. Insure the machine for $2000.00 and send your contact information. | They will come to your site, working you into their scheduled routes. It takes about 30 minutes. |
| Cost is about $90.00 + shipping | Cost is about $65.00 |

- The Pilot Audiometer Speech Frequencies – PASS = 25dB (30dB and above fail)

- The Pilot Audiometer Pure Tone -

    - 500 Hz - n/a

    - 1 kHz -  20dB = pass

    - 2 kHz - 20dB = pass

    - 4 kHz - 20dB = pass

- Tympanometer - Consult audiologist

Procedure for Vision Screenings

- Initial vision screenings will be done by Casey Eye Institute.

  - Contact Casey Eye Institute to schedule screening within 45 days of entry. Schedule ASAP. Casey Eye Institute Screening Program

- Center Managers will have primary responsibility for assuring that all children are screened within 45 days. If child's vision has not been screened within 30 days, a plan will be developed to involve the family to assure completion of screening before the 45 day deadline.

- All referrals will be completed by Winter Break for those children enrolling in September.

- If a child enters the program with eyeglasses, a vision screening should be entered into Child Plus and soaped in the child's file, that states “glasses” in the results, and a follow-up plan should be written to assure that child is getting annual evaluations by an optometrist.

- All returning children will receive a vision screening.

- All vision screening results and follow-up plans and activities will be documented in Child Plus using CP2b and in the child's file to communicate results and follow-up plans to families.

- For children enrolled throughout the year:

  - Contact Casey Eye Institute to schedule screening within 45 days of entry. Schedule ASAP. Casey Eye Institute Screening Program

  - When Casey Eye Institute is unavailable use the Preschool Vision Screening Kit

    - PASS = 4 or 5 correct on both tests

    - PASS – Distance Vision is 20/40

    - PASS – Random Dot E  if correctly indentifies E  4 out of 5 times

    - Not all centers have this screening tool and will need to work with Casey Eye Institute to ensure all children are screened within 45 days of entry.

Notes:

- Teachers help children become familiar with vision and hearing screening materials prior to screening through games incorporated into lessons or the environment.  This will help eliminate a false fail by children who may not be familiar with a picture or a vocabulary term.

- If the plan is to rescreen hearing before making a referral to the primary health provider, the rescreen should occur within 3-weeks or as soon as possible.  Base the decision on the individual child considering age, health and parental input.

- It is not necessary to screen children already screened or evaluated by an outside provider.  Get a release from the parent, verify and document the service on a CP2b.

This procedure complies with Head Start Performance Standards 45 CFR Section 1304.20.